| Customer | : ................................................................................... |  |
| --- | --- | --- |
| Software | : ICDAS 2.1 |  |
| Base Version No. | : 2.1 |  |
| Released Version No. | : 2.1.1 |  |
| System Location | : .................................................................................. |  |
| Protocol Pre Approval |  |  |
| Prepared By: | Name : ..................................................... Sign : ..................................................... Date : ..................................................... | Company ...................................................................... |
| Reviewed By: | Name : ..................................................... Sign : ..................................................... Date : ..................................................... | Company ...................................................................... |
| Approved By: | Name : ..................................................... Sign : ..................................................... Date : ..................................................... | Company ...................................................................... |
| Test Execution Details |  |  |
| Test Executed By: | Name : ..................................................... Sign : ..................................................... | Company ...................................................................... |
| Test Verified By: | Name : ..................................................... Sign : ..................................................... | Company ....................................................... |
| Test Date | Start of Test ................/...................../................... (dd/mmm/yyyy) | End of Test ................/...................../................ (dd/mmm/yyyy) |

Contents:

1	Protocol Pre approval	3

2	Brief Description	3

3	Purpose	3

4	Test procedures	3

4.1	Test rules	3

5	System Architecture	4

6	EQUIPMENT DISPLAYES	5

6.1	Equipment No 1	5

6.2	Equipment No 2	6

6.3	Equipment No 3	7

6.4	Equipment No 4	8

6.5	Equipment No 5	9

6.6	Equipment No 6	10

7	Verification of Software functionality	11

7.1	Contingency Testing	11

7.1.1	Alarms Contingency Testing	11

7.1.2	Network Failure Testing/Data logging Test	13

7.2	System Behaviour as per the USFDA 21 CFR PART 11 Clauses	15

8	Deviation Report	18

9	Appendix	19

10	Conclusion	20

11	Approval of test results	20

# Protocol Pre approval

The responsible person signing the Pre-Approval page (page no: 1) had reviewed and approved the Performance Qualification protocol for Newtronic ICDAS 2.1 Software. Signing of this protocol indicates that the contents of the documents had been reviewed, all test procedures are accurate and the acceptance criteria are applicable for the intended purpose of this study.

# Brief Description

The ICDAS 2.1 software qualification and documentation is carried at Newtronic Life care Equipment Pvt. Ltd. for functionality testing of all forms and features in ICDAS 2.1 software. This qualification document ensures the software is tested and qualified before release and can be used for future reference.

The ICDAS 2.1 is a browser based 21 CFR part 11 software used for Online Monitoring of Newtronic make Equipment’s, Data and Alarm Logging, Email Generation and transmission (Optional), SMS generation and Transmission (Optional).

# Purpose

The purpose of this document is to verify test for,

- Contingency Testing

- USFDA 21 CFR PART 11 Clauses Verification

# Test procedures

The performance of Software shall be verified by reviewing the installed systems and physical parameters, using the test data sheets provided in the protocol. The test data sheets will be used to document the performance of the software.

## Test rules

The PQ forms are intended for entering test notations pertaining to individual tests and any deviations that may occur.

All test result must be completed manually during the test execution, as well as for the remarks areas.

If the test is not applicable please write "NA" ( Not applicable). All corrections must be dated and done by hand with sign. The original information has to be crossed out by a single line. The corrected information must be clearly visible.

Each completed tests has to be signed and dated. Deviations from the PQ specification that occur must be registered in the deviation protocol and recorded in the appendix.

The PQ is concluded if all acceptance criteria are fulfilled and the notations / deviations registered in the deviation protocol are remedied or accepted.

# System Architecture

SYSTEM ACHITECTURE DIAGRAM

ICDAS 2.1 Lite Software

MSSQL

Centralized

DB

Browser Based Application

TCP IP Network

| Sr. No. | Technical data Version |  |  |  |  | Test result met (Yes/No) | Results and Comments |
| --- | --- | --- | --- | --- | --- | --- | --- |
|  | Description | Expected Result | Observation |  |  |  |  |
| # EQUIPMENT DISPLAYES |  |  |  |  |  |  |  |
| ## Equipment No 1 |  |  |  |  |  |  |  |
|  | Click On the Equipment Configuration option in the Configuration window | The Equipment Configuration Page Opens |  |  |  |  |  |
|  | Displays the Equipment details Equipment Name IP address Log Interval Door Access Type Click Save button | The required details should be available and displayed in the Equipment Configuration window | Details Entered Equipment Name:__________ IP Address:_______________ Log Interval:______________ Door Access Type:___________ Biometric Banner Text:_______________ Door Access User Name Status:________ Door Access User Details |  |  |  |  |
|  |  |  | User | User Name | User ID |  |  |
|  |  |  | Door Access1 |  |  |  |  |
|  |  |  | Door Access2 |  |  |  |  |
|  |  |  | Door Access3 |  |  |  |  |
|  |  |  | Door Access4 |  |  |  |  |
|  |  |  | Door Access5 |  |  |  |  |
|  |  |  | Door Access6 |  |  |  |  |

| Sr. No. | Technical data Version |  |  |  |  | Test result met (Yes/No) | Results and Comments |
| --- | --- | --- | --- | --- | --- | --- | --- |
|  | Description | Expected Result | Observation |  |  |  |  |
| ## Equipment No 2 |  |  |  |  |  |  |  |
|  | Click On the Equipment Configuration option in the Configuration window | The Equipment Configuration Page Opens |  |  |  |  |  |
|  | Displays the Equipment details Equipment Name IP address Log Interval Door Access Type Click Save button | The required details should be available and displayed in the Equipment Configuration window | Details Entered Equipment Name:__________ IP Address:_______________ Log Interval:______________ Door Access Type:___________ Biometric Banner Text:_______________ Door Access User Name Status:________ Door Access User Details |  |  |  |  |
|  |  |  | User | User Name | User ID |  |  |
|  |  |  | Door Access1 |  |  |  |  |
|  |  |  | Door Access2 |  |  |  |  |
|  |  |  | Door Access3 |  |  |  |  |
|  |  |  | Door Access4 |  |  |  |  |
|  |  |  | Door Access5 |  |  |  |  |
|  |  |  | Door Access6 |  |  |  |  |

| Sr. No. | Technical data Version |  |  |  |  | Test result met (Yes/No) | Results and Comments |
| --- | --- | --- | --- | --- | --- | --- | --- |
|  | Description | Expected Result | Observation |  |  |  |  |
| ## Equipment No 3 |  |  |  |  |  |  |  |
|  | Click On the Equipment Configuration option in the Configuration window | The Equipment Configuration Page Opens |  |  |  |  |  |
|  | Displays the Equipment details Equipment Name IP address Log Interval Door Access Type Click Save button | The required details should be available and displayed in the Equipment Configuration window | Details Entered Equipment Name:__________ IP Address:_______________ Log Interval:______________ Door Access Type:___________ Biometric Banner Text:_______________ Door Access User Name Status:________ Door Access User Details |  |  |  |  |
|  |  |  | User | User Name | User ID |  |  |
|  |  |  | Door Access1 |  |  |  |  |
|  |  |  | Door Access2 |  |  |  |  |
|  |  |  | Door Access3 |  |  |  |  |
|  |  |  | Door Access4 |  |  |  |  |
|  |  |  | Door Access5 |  |  |  |  |
|  |  |  | Door Access6 |  |  |  |  |

| Sr. No. | Technical data Version |  |  |  |  | Test result met (Yes/No) | Results and Comments |
| --- | --- | --- | --- | --- | --- | --- | --- |
|  | Description | Expected Result | Observation |  |  |  |  |
| ## Equipment No 4 |  |  |  |  |  |  |  |
|  | Click On the Equipment Configuration option in the Configuration window | The Equipment Configuration Page Opens |  |  |  |  |  |
|  | Displays the Equipment details Equipment Name IP address Log Interval Door Access Type Click Save button | The required details should be available and displayed in the Equipment Configuration window | Details Entered Equipment Name:__________ IP Address:_______________ Log Interval:______________ Door Access Type:___________ Biometric Banner Text:_______________ Door Access User Name Status:________ Door Access User Details |  |  |  |  |
|  |  |  | User | User Name | User ID |  |  |
|  |  |  | Door Access1 |  |  |  |  |
|  |  |  | Door Access2 |  |  |  |  |
|  |  |  | Door Access3 |  |  |  |  |
|  |  |  | Door Access4 |  |  |  |  |
|  |  |  | Door Access5 |  |  |  |  |
|  |  |  | Door Access6 |  |  |  |  |

| Sr. No. | Technical data Version |  |  |  |  | Test result met (Yes/No) | Results and Comments |
| --- | --- | --- | --- | --- | --- | --- | --- |
|  | Description | Expected Result | Observation |  |  |  |  |
| ## Equipment No 5 |  |  |  |  |  |  |  |
|  | Click On the Equipment Configuration option in the Configuration window | The Equipment Configuration Page Opens |  |  |  |  |  |
|  | Displays the Equipment details Equipment Name IP address Log Interval Door Access Type Click Save button | The required details should be available and displayed in the Equipment Configuration window | Details Entered Equipment Name:__________ IP Address:_______________ Log Interval:______________ Door Access Type:___________ Biometric Banner Text:_______________ Door Access User Name Status:________ Door Access User Details |  |  |  |  |
|  |  |  | User | User Name | User ID |  |  |
|  |  |  | Door Access1 |  |  |  |  |
|  |  |  | Door Access2 |  |  |  |  |
|  |  |  | Door Access3 |  |  |  |  |
|  |  |  | Door Access4 |  |  |  |  |
|  |  |  | Door Access5 |  |  |  |  |
|  |  |  | Door Access6 |  |  |  |  |

| Sr. No. | Technical data Version |  |  |  |  | Test result met (Yes/No) | Results and Comments |
| --- | --- | --- | --- | --- | --- | --- | --- |
|  | Description | Expected Result | Observation |  |  |  |  |
| ## Equipment No 6 |  |  |  |  |  |  |  |
|  | Click On the Equipment Configuration option in the Configuration window | The Equipment Configuration Page Opens |  |  |  |  |  |
|  | Displays the Equipment details Equipment Name IP address Log Interval Door Access Type Click Save button | The required details should be available and displayed in the Equipment Configuration window | Details Entered Equipment Name:__________ IP Address:_______________ Log Interval:______________ Door Access Type:___________ Biometric Banner Text:_______________ Door Access User Name Status:________ Door Access User Details |  |  |  |  |
|  |  |  | User | User Name | User ID |  |  |
|  |  |  | Door Access1 |  |  |  |  |
|  |  |  | Door Access2 |  |  |  |  |
|  |  |  | Door Access3 |  |  |  |  |
|  |  |  | Door Access4 |  |  |  |  |
|  |  |  | Door Access5 |  |  |  |  |
|  |  |  | Door Access6 |  |  |  |  |

| Sr. No. | Technical data Version |  |  |  | Test result met (Yes/No) | Results and Comments |
| --- | --- | --- | --- | --- | --- | --- |
|  | Description | Expected Result | Observation |  |  |  |
| # Verification of Software functionality |  |  |  |  |  |  |
| ## Contingency Testing |  |  |  |  |  |  |
| ### Alarms Contingency Testing |  |  |  |  |  |  |
|  | Description | The ICDAS 2.1 Lite software must be able to download stored readings in the memory of PLC to ICDAS 2.1 Lite software of every connected equipment |  |  |  |  |
|  | Test Procedure | Generate the alarms as mentioned below and check whether it gets transffered to the software |  |  |  |  |
|  | Does the logged alarms from Equipment ____________ gets transferred to software | The logged alarms in PLC must get transferred to software |  |  |  |  |
|  | Does the logged alarms from Equipment ____________ gets transferred to software | The logged alarms in PLC must get transferred to software |  |  |  |  |
|  | Does the logged alarms from Equipment ____________ gets transferred to software | The logged alarms in PLC must get transferred to software |  |  |  |  |
|  | Does the logged alarms from Equipment ____________ gets transferred to software | The logged alarms in PLC must get transferred to software |  |  |  |  |
|  | Does the logged alarms from Equipment ____________ gets transferred to software | The logged alarms in PLC must get transferred to software |  |  |  |  |
|  | Does the logged alarms from Equipment ____________ gets transferred to software | The logged alarms in PLC must get transferred to software |  |  |  |  |

| Sr. No. | Technical data Version |  |  |  | Test result met (Yes/No) | Results and Comments |
| --- | --- | --- | --- | --- | --- | --- |
|  | Description | Expected Result | Observation |  |  |  |
| ### Network Failure Testing/Data logging Test |  |  |  |  |  |  |
|  | Description | The ICDAS 2.1 Lite software must be able to download stored readings and alarms in the memory of PLC to ICDAS 2.1 Lite software of every connected equipment |  |  |  |  |
|  | Test Procedure | - Shut down the Local Client PC for 30 mins OR Stop ICDAS Local Service. - Generate Alarms for Door Open/Close, Mains Fail/Resume & Water Level Fail/Ok - Restart the Client PC and cross check whether the data and alarms is downloaded or not. |  |  |  |  |
| 1. | Does the data and alarms stored from Equipment ID ____________ gets transferred to software | The stored reading and alarms must get transferred to the ICDAS 2.1 software once network is connected or ok |  |  |  |  |
| 2. | Does the data and alarms stored from Equipment ID ____________ gets transferred to software | The stored reading and alarms must get transferred to the ICDAS 2.1 software once network is connected or ok |  |  |  |  |
| 3. | Does the data and alarms stored from Equipment ID ____________ gets transferred to software | The stored reading and alarms must get transferred to the ICDAS 2.1 software once network is connected or ok |  |  |  |  |
| 4. | Does the data and alarms stored from Equipment ID ____________ gets transferred to software | The stored reading and alarms must get transferred to the ICDAS 2.1 software once network is connected or ok |  |  |  |  |
| 5. | Does the data and alarms stored from Equipment ID ____________ gets transferred to software | The stored reading and alarms must get transferred to the ICDAS 2.1 software once network is connected or ok |  |  |  |  |
| 6. | Does the data and alarms stored from Equipment ID ____________ gets transferred to software | The stored reading and alarms must get transferred to the ICDAS 2.1 software once network is connected or ok |  |  |  |  |
| 7. | Reports | The software must print the records in PDF and with highlights(Bold) in case of deviation |  |  |  |  |

| Sr. No. | Sr. No. (As per Clauses) | Questions | Points to be verified | Observation | Accepted? (Y/N/NA) | Reference No | Results and Comments |
| --- | --- | --- | --- | --- | --- | --- | --- |
| ## System Behaviour as per the USFDA 21 CFR PART 11 Clauses |  |  |  |  |  |  |  |
|  | 11.10 (a) | Is it possible to discern invalid or altered records? | If data logging facilities available then verify the possibilities of the record alteration. |  |  |  |  |
|  | 11.10 (b) | Is the system capable of producing accurate and complete copies of electronic records on paper? | Verify the reproducibility of historical data report |  |  |  |  |
|  | 11.10 (b) | Is the system capable producing accurate and complete copies of records in electronic form for inspection, review, and copying by the FDA? | Verify facilities of data backup if data is being stored in electronic form by system |  |  |  |  |
|  | 11.10 (c) | Are the records readily retrievable throughout their retention period? | Verify data backup location and retrievable facilities. |  |  |  |  |
|  | 11.10 (d) | Is there authorized access available for individual login? | Verify the authorized access to software. |  |  |  |  |
|  | 11.10 (e) | Is there a secure, computer generated, time stamped audit trail that records the date and time of operator entries and actions that create, modify, or delete electronic records? | Verify that audit trails are available for any database related activities. |  |  |  |  |
|  | 11.10 (e) | Upon making a change to an electronic record, is previously recorded information still available (i.e. Not obscured by the change)? | Verify that in case of modification or deletion of any stored data, old data should be available. |  |  |  |  |
|  | 11.10 (e) | Is an electronic record’s audit trail retrievable throughout the record’s retention period? | Verify that audit trail is available till data retention period. |  |  |  |  |
|  | 11.10 (e) | Is the audit trail available for review and copying by the FDA? | Verify facilities of audit trail backup and view. |  |  |  |  |
|  | 11.10 (f) | If the sequence of system steps or events is important, is this enforced by the system (e.g. As would be the case in a process control system)? | Verify critical event and alarm data logging facilities |  |  |  |  |
|  | 11.10 (g) | Is there access level for individual available for authorized access and perform operation? | Verify for access level availability |  |  |  |  |
|  | 11.10 (h) | If it is a requirement of the system that input data or instructions can only come from certain input devices (e.g. Terminals) does the system check the validity of the source of any data or instructions received? (Note: this applies where data or instructions can come from more than one device, and therefore the system must verify the integrity of its source). | Verify system compatibilities between different devices (mainly communication network and software related) |  |  |  |  |
|  | 11.10 (i) | Is there documented training, including on the job training for system users, developers, it support staff? | Verify the training record. Document the responsibility for all system user and related responsible person. |  |  |  |  |
|  | 11.10 (j) | Is there a written policy that makes individuals fully accountable and responsible for actions initiated under their electronic signature? | Verify the written procedure for electronic signatures. |  |  |  |  |
|  | 11.10 (k) | Is the distribution of, access to, and use of system operation and maintenance documentation controlled? | Same as 11.10(i) & (j) |  |  |  |  |

# Deviation Report

Report Number: ………………………………………………….

Deviations to the acceptance criteria that occurred during testing need to be stated in the deviation report. One deviation report must be created for each deviation and the report number to be indicated accordingly.

The deviation report can be copied as required.

| Description of problem | ……………………………………………………………………………………… ……………………………………………………………………………………… ……………………………………………………………………………………… ……………………………………………………………………………………… ……………………………………………………………………………………… ……………………………………………………………………………………… |
| --- | --- |
| Measures for trouble shooting | ……………………………………………………………………………………… ……………………………………………………………………………………… ……………………………………………………………………………………… ……………………………………………………………………………………… ……………………………………………………………………………………… ……………………………………………………………………………………… ……………………………………………………………………………………… |
| Remarks | ……………………………………………………………………………………… ……………………………………………………………………………………… ……………………………………………………………………………………… |
| Executed By Verified/ accepted by | Name:………………………Date: . . . . . . . . . . . . . . . . Signature: . . . . .. . . . . . . . . . . . . . . Name:………………………Date: . . . . . . . . . . . . . . . . Signature: . . . . .. . . . . . . . . . . . . . |

# Appendix

| Annexure Number | Contents | Number of pages |
| --- | --- | --- |
| PQ_ICDAS2.1_Lite/ |  |  |
| PQ_ICDAS2.1_Lite/ |  |  |
| PQ_ICDAS2.1_Lite/ |  |  |
| PQ_ICDAS2.1_Lite/ |  |  |

# Conclusion

. . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . .  . . . . . . . . . . . . . . . . . . . . . . . . . .

. . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . .. . . . . . . . . . . . . . . . . . . . . . . . . . . .

. . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . .  . . . . . . . . . . . . . . . . . . . . . . . . . . .

. . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . .  . . . . . . . . . . . . . . . . . . . . . . . . . . .

. . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . .. . . . . . . . . . . . . . . . . . . . . . . . . . . .

. . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . .  . . . . . . . . . . . . . . . . . . . . . . . . . . .

. . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . .. . . . . . . . . . . . . . . . . . . . . . . . . . . .

. . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . .. . . . . . . . . . . . . . . . . . . . . . . . . . . .

. . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . .. . . . . . . . . . . . . . . . . . . . . . . . . . . .

. . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . .. . . . . . . . . . . . . . . . . . . . . . . . . . . .

. . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . .. . . . . . . . . . . . . . . . . . . . . . . . . . . . .

# Approval of test results

The Performance Qualification is approved if all acceptance criteria are fulfilled and the notations / deficiencies registered in the deviation protocol are corrected or fully accepted.

Approved By:

| Tests executed | Name | Signature and Date |
| --- | --- | --- |
| Executed By | . . . . . . . . . . . . . . . . . . . | . . . . . . . . . . . . . . . . . . |
| All Test criteria have been fulfilled | Name | Signature and Date |
| Verified By | . . . . . . . . . . . . . . . . . . . | . . . . . . . . . . . . . . . . . . |
| All Test criteria have been fulfilled | Name | Signature and Date |
| Approved By | . . . . . . . . . . . . . . . . . . . | . . . . . . . . . . . . . . . . . . |