| Customer |  | ............................................................................................................................ |  |  |
| --- | --- | --- | --- | --- |
| Software |  | Neclog 3.0 |  |  |
| Version No. |  | 3.0 |  |  |
| System Location |  | ............................................................................................................................ |  |  |
| Protocol Pre Approval |  |  |  |  |
| Prepared By: |  |  | Name : ................................................................. Sign : ................................................................. Date : ................................................................. | Company ...................................................................................... |
| Reviewed By: |  |  | Name : ................................................................. Sign : ................................................................. Date : ................................................................. | Company .................................................................................... |
| Approved By: |  |  | Name : ................................................................. Sign : ................................................................. Date : ................................................................. | Company .................................................................................... |
| Test Execution Details |  |  |  |  |
| Test Executed By: |  |  | Name : ................................................................. Sign : ................................................................. | Company ................................................................................... |
| Test Verified By: |  |  | Name : ................................................................ Sign : ................................................................ | Company ................................................................................... |
| Test Date |  |  | Start of Test ................/...................../................... (dd/mmm/yyyy) | End of Test ................/...................../................ (dd/mmm/yyyy) |

Contents:

1	Protocol Pre approval	5

2	Brief Description	5

3	Purpose	5

4	Test procedures	5

4.1	Test rules	5

5	System Architecture	6

5.1	Server Architecture	6

6	Verification of Software functionality	7

6.1	GUI (Graphical User Interface) Function Screen Testing	7

7	Instrument Addition	8

8	GUI Screens	9

8.1	Home	9

8.2	General Symbols	9

8.2.1	Change Password	10

8.2.2	Switch Department	10

9	Administration	11

9.1	Company	11

9.2	Activation Details	12

9.3	Department	12

9.4	Application Setup	14

9.5	User Group	15

9.6	Admin User	16

9.7	Backup	17

9.8	System Diagnosis	17

9.9	Data Archive	17

10	Configuration	18

10.1	Application User	18

10.2	Unit Master	19

10.3	Instrument Config	19

10.4	Synchronize RTC	20

10.5	Email/SMS Alert	21

10.6	Import Files	22

11	Real Time	22

11.1	Real Time Display	22

12	Data Log	23

12.1	Manage Data Log	23

12.1.1	Review Data Log	23

12.1.2	Approve Data Log	24

12.2	View Data Log	24

12.3	View Reviewed Data	25

12.4	View Approved Data	25

12.5	Generate MKT	26

12.6	MKT Analysis	26

12.7	Alarm Log	27

13	Audit Trail	28

13.1	Alarm Audit Trail	28

13.2	Instrument Audit Trail	28

13.3	User Audit Trail	29

13.4	Email Audit Trails (Only if Email Alert Is Active)	30

13.5	SMS Audit Trails (If SMS is Active)	30

13.6	Reviewed/ Approved Audit Trails	31

14	Archive Analysis	32

14.1	Archive Analysis: -	32

14.1.1	Database Restore	32

14.2	Data Log Analysis	35

14.2.1	View Data Log	35

14.2.2	View Reviewed Log	35

14.2.3	View Approved Data	36

14.2.4	MKT Analysis	36

14.2.5	Alarm Log	37

14.3	Audit Trail	37

14.4	Data Restore	38

15	Company Information	38

16	User Access Rights Groups Verification	39

17	Password Integrity---Admin User	40

17.1	Login name Uniqueness / Password Complexity	40

17.2	Password Integrity	40

18	Password Integrity—Application User	41

18.1	Login name Uniqueness / Password Complexity	41

18.2	Password Integrity	41

19	Admin User Addition / Deactivation	42

20	Application User Addition / Deactivation	43

21	Admin User Access Rights Verification	44

22	Application User Access Rights Verification	45

23	Software Functionality testing	46

23.1	Power Failure Condition	46

23.2	Auto Backup	46

23.3	Auto Logout	46

23.4	Auto Password Save	46

24	Audit trail Functionality testing	47

25	References / Glossary	48

26	Deviation Report	49

27	Appendix	50

28	Conclusion	53

29	Approval of test results	53

# Protocol Pre approval

The responsible person signing the Pre-Approval page (page no: 1) had reviewed and approved the     Operational Qualification protocol for Newtronic Neclog 3.0 Software. Signing of this protocol indicates that the contents of the documents had been reviewed, all test procedures are accurate and the acceptance criteria are applicable for the intended purpose of this study.

# Brief Description

The Neclog 3.0 software qualification and documentation is carried at Newtronic Lifecare Equipment Pvt. Ltd. for functionality testing of all forms and features in Neclog 3.0 software. This qualification document ensures the software is tested and qualified before release and can be used for future reference.

The Neclog 3.0 is a browser based 21 CFR part 11 software used for logging of parameters temperature and humidity from the Newtronic data logger NW70. The software is a browser based application and is 21 CFR ready software with all features available.

# Purpose

The purpose of this document is to verify test for,

- Forms, GUI Screens function and menu bar items

- System Securities/Password Integrity

- Power Failure Availability of required utilities.

- Alarms verification

- System Audit Trails.

# Test procedures

The operation of Software shall be verified by reviewing the installed systems and physical parameters, using the test data sheets provided in the protocol. The test data sheets will be used to document the installation of the software.

## Test rules

The OQ forms are intended for entering test notations pertaining to individual tests and any deviations that may occur.

All empty test result must be completed manually during the test execution, as well as for the remarks areas.

If the test is not applicable please marked "N / A” (not applicable) or "N / R" (nothing to report). All corrections must be dated and done by hand. The original info has to be crossed out by a single line. The corrected information must be clearly visible.

Each completed tests has to be signed and dated. Deviations from the OQ specification that occur must be registered in the deviation protocol and recorded in the appendix.

The OQ is concluded if all acceptance criteria are fulfilled and the notations / deviations registered in the deviation protocol are remedied or accepted.

# System Architecture

## Server Architecture

USB

USB

FTP

Centralised Server

| Sr. No. | Technical data Version |  |  | Observation | Test result met (Yes/No) |  | Results and Comments |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- |
| # Verification of Software functionality |  |  |  |  |  |  |  |  |
|  | ## GUI (Graphical User Interface) Function Screen Testing |  |  |  |  |  |  |  |
|  | Description | Describes the details of functionality of Forms/GUI screens and operations of main menu items of main screens/forms that is the task bar pane screens or more popularly called as the drop down menus. |  |  |  |  |  |  |
|  | Test Procedure | - Login into the Neclog 3.0 Software with correct security level to enable all functions – here Administrator login enables all functions. - Expand the folders in the Main Menu window/explorer-panel, select the desired form and double-click on its icon to open its window. - Expand each & every folder and subfolders and open the forms. Verify & record the results. |  |  |  |  |  |  |
| - Common Buttons |  |  |  |  |  |  |  |  |
|  | Login |  |  |  |  |  |  |  |
|  | Add |  |  |  |  |  |  |  |
|  | Save |  |  |  |  |  |  |  |
|  | Edit |  |  |  |  |  |  |  |
|  | Update |  |  |  |  |  |  |  |
|  | Delete |  |  |  |  |  |  |  |
|  | Print |  |  |  |  |  |  |  |
|  | Find |  |  |  |  |  |  |  |

| Sr. No. | Technical data Version |  | Observation | Test result met (Yes/No) | Results and Comments |
| --- | --- | --- | --- | --- | --- |
|  | # Instrument Addition |  |  |  |  |
|  | Description | To test all qualification test cases in software we need to add one instrument in it. |  |  |  |
|  | Test Procedure | - Login into the Neclog 3.0 Software with correct security level to enable all functions – here Administrator login enables all functions. - Expand the folders in the Main Menu window/explorer-panels, select the configuration menu, and go to instruments configuration double-click on its icon to open its window. |  |  |  |
|  | Click On the Instrument Configuration option in the Configuration window | The Instrument Configuration Page Opens |  |  |  |
|  | Click On ADD to add a new Instrument | The Add Instrument window must open for entering of new instrument |  |  |  |
|  | Enter the following Details 1.Instrument Name 2.Type 3.IP address 4.Serial Number 5.Log Interval 6.Active Probes 7.Mode 8.Status Click Save & continue | The Add Instrument window must open for entering of new Instrument’s details | Details Entered 1.Instrument Name :- 2.Type :- 3.Instrument IP :- 4.Serial Number :- 5.Log Interval :- 6.Active Probes :- 7.Mode :- 8.Status :- |  |  |
|  | Password window with comment must open | password with comment window open |  |  |  |

| Sr. No. |  |  | Technical data Version |  | Observation | Test result met (Yes/No) |  | Results and Comments |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| # GUI Screens |  |  |  |  |  |  |  |  |  |
| ## Home |  |  |  |  |  |  |  |  |  |
|  | Login to Software and click on Home option |  |  | Home screen must open and must display the summary of Instrument’s and unattended data/alarms which instruments is Online in software. |  |  |  |  |  |
|  | List View |  |  | Must display the summary of online instrument’s in tabular listed form |  |  |  |  |  |
| ## General Symbols |  |  |  |  |  |  |  |  |  |
|  |  | After clicking on symbol, it is shows change password & Logout window option. |  | Must display change password & logout window. |  |  |  |  |  |

| Sr. No. | Technical data Version |  | Observation | Test result met (Yes/No) | Results and Comments |
| --- | --- | --- | --- | --- | --- |
| ### Change Password |  |  |  |  |  |
|  | Click On the Change Password | Change Password should Opens |  |  |  |
|  | Enter the following Details 1.OLD Password 2.New Password 3.Confirm New Password Click Change | Alpha Numeric Minimum ******** and max.*************** code Does the password get updated to new password when click on change? |  |  |  |
|  | After changing software get auto logout | Software must auto Logout after changing the password |  |  |  |
|  | Click on login to cross check the password change | Entering login name and new password must give access to software? |  |  |  |
| ### Switch Department |  |  |  |  |  |
| a) | Click on department | Department switch window open to switch departments. |  |  |  |
| a) | Select department | Click on select Department. |  |  |  |

| Sr. No. | Technical data Version |  | Observation | Test result met (Yes/No) | Results and Comments |
| --- | --- | --- | --- | --- | --- |
| # Administration 9.1 Company 9.6 Admin User 9.2 Activation Key 9.7 Backup 9.3 Department 9.8 System Diagnosis 9.4 Application Setup 9.9 Archive Data 9.5 User Group |  |  |  |  |  |
| ## Company |  |  |  |  |  |
|  | Click on Company option in Administrator window | The company Details Info window will open and display - Company Name - Address - Image |  |  |  |
|  | Company Name | Displays/edits the entered Company details |  |  |  |
|  | Address | Displays/edits the entered address |  |  |  |
|  | Image | Attach of company logo mage of specified size |  |  |  |
|  | Update Button | The Available buttons on the window must function as specified in the 6.1 |  |  |  |

| Sr. No. | Technical data Version |  | Observation | Test result met (Yes/No) | Results and Comments |
| --- | --- | --- | --- | --- | --- |
| ## Activation Details |  |  |  |  |  |
|  | To Activate new licence click Activation Details Enter the following Details - Activation Key - Click on Update | Enter Activation Key & Update it. |  |  |  |
|  | Buttons | The Available buttons on the window must function as specified in the 6.1 |  |  |  |
| ## Department |  |  |  |  |  |
|  | Click on Department option in Administrator window | The Department window will open and must display the added departments |  |  |  |
|  | To add/edit new department click on add/edit option and Enter the following Department Details Department Name Header Note Footer Note Archive Month | Details Entered/edited |  |  |  |

| Sr. No. | Technical data Version |  |  |  | Observation | Test result met (Yes/No) |  |  | Results and Comments |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
|  | Email Email System Email Alert Delay Email Time Email Address 1 Email Address 2 Email Address 3 Email Address 4 Email Address 5 Supervisor Email Add1 Supervisor Email Add2 |  | Details Entered/ edited |  |  |  |  |  |  |  |
|  | SMS SMS System SMS Alert Delay Mobile User Name 1 Mobile User Name 2 Mobile User Name 3 Mobile User Name 4 Mobile User Name 5 Mobile Number 1 Mobile Number 2 Mobile Number 3 Mobile Number 4 Mobile Number 5 Click on Save |  | Details Entered/ edited |  |  |  |  |  |  |  |
|  | Buttons |  | The Available buttons on the window must function as specified in the 6.1 |  |  |  |  |  |  |  |
| ## Application Setup |  |  |  |  |  |  |  |  |  |  |
|  |  | Click on Application Setup option in Administrator window |  | The Application setup window will open and display the details |  |  |  |  |  |  |
|  |  | The Allowable Fields can be added/edited Import File Path Backup local Path Backup remote Path Backup Time Report Date/Time Stamp Password Change System Auto logout User Access Lock Count Email System SMTP Authentication SMTP Email SMTP Address SMTP User Name SMTP Password SMTP Port IMAP Authentication IMAP Email IMAP Password IMAP Port Secure Sockets Layer Email Signature SMS System Communication Port Parity Baud Rate Data Bits Stops Bits Click on Update |  | Details Entered/edited |  |  |  |  |  |  |
| Sr. No. |  | Technical data Version |  |  | Observation | Test result met (Yes/No) |  | Results and Comments |  |  |
|  |  | Test Email Via SMPT |  | Field to check email |  |  |  |  |  |  |
|  |  | Test Email Via IMAP |  | Field to check email |  |  |  |  |  |  |
|  |  | SMS Test |  | Field to check SMS |  |  |  |  |  |  |
|  |  | Buttons |  | The Available buttons on the window must function as specified in the 6.1 |  |  |  |  |  |  |
| ## User Group |  |  |  |  |  |  |  |  |  |  |
|  | Click on User Group option in Administrator window |  |  | The User Group window will open and must display the added user groups |  |  |  |  |  |  |
|  | To add/edit new User group click on Add/edit and enter User Group Details Select Access Rights Click on Save |  |  | Details Entered/edited |  |  |  |  |  |  |
|  | Buttons |  |  | The Available buttons on the window must function as specified in the 6.1 |  |  |  |  |  |  |

| Sr. No. | Technical data Version |  | Observation | Test result met (Yes/No) | Results and Comments |
| --- | --- | --- | --- | --- | --- |
| ## Admin User |  |  |  |  |  |
|  | Click on Admin User option in Administrator window | The Admin User window will open and must display the added admin user |  |  |  |
|  | To add/edit admin user Enter the following Details User Name Login Name Password Department Name Accessible Department User Group Set Password Duration Status Click on Save | Details Entered/edited |  |  |  |
|  | Buttons | The Available buttons on the window must function as specified in the 6.1 |  |  |  |

| Sr. No. | Technical data Version |  | Observation | Test result met (Yes/No) | Results and Comments |
| --- | --- | --- | --- | --- | --- |
| ## Backup |  |  |  |  |  |
|  | Click on Backup option in Administrator window Click on Backup Data | The Backup must happen at Path in Application Settings |  |  |  |
| ## System Diagnosis |  |  |  |  |  |
|  | Instrument Log | Displays the Status of instruments |  |  |  |
|  | Error Log | Displays the Error log while transferring the logs from Instruments |  |  |  |
|  | Restart Service | Restart Service |  |  |  |
| ## Data Archive |  |  |  |  |  |
|  | Department Name | Department data to be archived |  |  |  |
|  | Data Archive Date Click Archive | Date till which the data gets archived if available |  |  |  |

| Sr. No. | Technical data Version |  | Observation | Test result met (Yes/No) | Results and Comments |
| --- | --- | --- | --- | --- | --- |
| # Configuration 10.1 Application User 10.4 Synchronize RTC 10.2 Unit Master 10.5 Email/ SMS Alert 10.3 Instrument config 10.6 Import Files |  |  |  |  |  |
| ## Application User |  |  |  |  |  |
|  | Click on the Application User option in the Configuration Window | The Application User window Opens |  |  |  |
|  | Click On add/edit button and enter the details User Name Login Name Password Department Name User Group Set Password Duration Status Click On Save | Opens the Login Details window for authentication |  |  |  |
|  | Buttons | The Available buttons on the window must function as specified in the 6.1 |  |  |  |

| Sr. No. | Technical data Version |  |  |  | Observation |  | Test result met (Yes/No) | Results and Comments |
| --- | --- | --- | --- | --- | --- | --- | --- | --- |
| ## Unit Master |  |  |  |  |  |  |  |  |
| a) | Click On the unit master option in the configuration | The Unit Master window will open and must display the added units |  |  |  |  |  |  |
| b) | Click on add/edit button and enter the details Unit Name Unit Type Click On Save | The Details entered should gets saved |  |  |  |  |  |  |
| c) | Buttons | The Available buttons on the window must function as specified in the 6.1 |  |  |  |  |  |  |
| ## Instrument Config |  |  |  |  |  |  |  |  |
|  | Click On the Instrument Config option in the Configuration |  | The Instrument Configuration Page Opens |  |  |  |  |  |
|  | Click On add/edit button and enter the details Instrument Name Type Instrument IP Serial No. Log Interval Active Probes Mode Status Click On Save & Continue |  | The details entered gets saved |  | Details Edited Instrument |  |  |  |

| Sr. No. | Technical data Version |  |  | Observation | Test result met (Yes/No) | Results and Comments |
| --- | --- | --- | --- | --- | --- | --- |
|  | Click On Temperature/ humidity / Pressure Probes Name Unit Resolution Graph Colour Lower Limit Upper Limit Axis Click On Save & Continue | The details entered & Login window open after putting password & Comment the changes gets saved |  | Details Edited Instrument |  |  |
|  | Buttons | The Available buttons on the window must function as specified in the 6.1 |  |  |  |  |
| ## Synchronize RTC |  |  |  |  |  |  |
|  | Click on the Synchronize RTC option in the Configuration |  | The Synchronize RTC window Opens and displays Instrument details |  |  |  |
|  | Select the Instrument Click on Synchronize |  | Opens the Login Details window for authentication |  |  |  |
|  | Software will ask for confirmation and Login Details Click On Login |  | Enter valid login details & Comment Will update the RTC of the Instrument with that of computer |  |  |  |
|  | Buttons |  | The Available buttons on the window must function as specified in the 6.1 |  |  |  |

| Sr. No. | Technical data Version |  | Observation | Test result met (Yes/No) | Results and Comments |
| --- | --- | --- | --- | --- | --- |
| ## Email/SMS Alert |  |  |  |  |  |
|  | Click on the Email/SMS alert option in the Configuration Window | The Email/SMS window Opens and displays added Instrument details |  |  |  |
|  | Select the Instrument Click on Edit | Opens the Alarms setting window for selection / deselection of alarms for Email/SMS |  |  |  |
|  | Select/Deselect the Alarms for sending via SMS/Email Click on Save | The Email/SMS must be sent only for the selected alarm events |  |  |  |
|  | Buttons | The Available buttons on the window must function as specified in the 6.1 |  |  |  |

| Sr. No. | Technical data Version |  | Observation | Test result met (Yes/No) | Results and Comments |
| --- | --- | --- | --- | --- | --- |
| ## Import Files |  |  |  |  |  |
|  | Click on the Import Files in the Configuration Window | The Import Files page Opens |  |  |  |
|  | Click On Import files and select the details Instrument File Type Select files Click On Upload files | The Import files must Upload in the software |  |  |  |
|  | Buttons | The Available buttons on the window must function as specified in the 6.1 |  |  |  |
| # Real Time |  |  |  |  |  |
| ## Real Time Display |  |  |  |  |  |
| a) | Click on the real time display option in the real time | The Real Time Display window will opens |  |  |  |
| b) | Real Time Display Window | Display Graph along with PV of the all Sensors |  |  |  |
| c) | Reset | Reset the Graph |  |  |  |
| d) | Buttons | The Available buttons on the window must function as specified in the 6.1 |  |  |  |

| Sr. No. | Technical data Version |  | Observation | Test result met (Yes/No) | Results and Comments |
| --- | --- | --- | --- | --- | --- |
| # Data Log 12.1 Manage Data Log 12.5 MKT Analysis 12.2 View data log 12.6 Generate MKT 12.3 View reviewed data 12.7 Alarm Log 12.4 View approved data |  |  |  |  |  |
| ## Manage Data Log |  |  |  |  |  |
| ### Review Data Log |  |  |  |  |  |
|  | Click On Review Data Log in the Data Log | The Review Data Log window will open |  |  |  |
|  | Click On Review | Print preview of filtered data report opens in PDF format Close PDF report then authentication window must open |  |  |  |
|  | Authentication window Open | After login the authentication data gets transfer in View Reviewed Data |  |  |  |
|  | Check the View Reviewed Data log for reviewed data | The reviewed data must get displayed in the View Reviewed data log |  |  |  |

| Sr. No. | Technical data Version |  |  | Observation | Test result met (Yes/No) | Results and Comments |
| --- | --- | --- | --- | --- | --- | --- |
| ### Approve Data Log |  |  |  |  |  |  |
|  | Click On Approve Data Log in the Data Log option | The Approve Data Log window will open |  |  |  |  |
|  | Click On Approve | Print preview of filtered data opens in PDF format Close PDF Report then authentication window must open |  |  |  |  |
|  | Authentication window Open | After login the authentication data gets transfer in View Approved Data |  |  |  |  |
|  | Check the View Approved Data log for Approved data | The reviewed data must get displayed in the View Approved data log |  |  |  |  |
| ## View Data Log |  |  |  |  |  |  |
|  | Click On the View Data Log in Data Log option |  | The View Data log window will open |  |  |  |
|  | Enter the Filter Option Click Print |  | Print Preview of filtered data in PDF format |  |  |  |
|  | Select Data log Graph |  | The Graph for the selected filtered details must get displayed |  |  |  |

| Sr. No. |  | Technical data Version |  |  |  |  |  | Observation | Test result met (Yes/No) | Results and Comments |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
|  |  | Buttons |  |  |  | The Available buttons on the window must function as specified in the 6.1.1 |  |  |  |  |
| ## View Reviewed Data |  |  |  |  |  |  |  |  |  |  |
|  |  | Click On the View Reviewed Data Log option |  |  |  | The View Reviewed Data log window will open |  |  |  |  |
|  |  | Enter the Filter Option Click on Print |  |  |  | Print preview of filtered data report opens in PDF format |  |  |  |  |
|  |  | Click On Graph |  |  |  | The Graph for the selected filtered details must get displayed |  |  |  |  |
|  |  | Buttons |  |  |  | The Available buttons on the window must function as specified in the 6.1.1 |  |  |  |  |
| ## View Approved Data |  |  |  |  |  |  |  |  |  |  |
|  |  | Click On the View Approved Data Log in Data Log Analysis |  |  | The View Approved Data log window will open |  |  |  |  |  |
|  |  | Enter the Filter Option Click Print |  |  | Print Preview of filtered data report opens in PDF format |  |  |  |  |  |
| Sr. No. |  | Technical data Version |  |  |  |  |  | Observation | Test result met (Yes/No) | Results and Comments |
|  |  | Click On Graph |  |  | The Graph for the selected filtered details must get displayed |  |  |  |  |  |
|  |  | Buttons |  |  | The Available buttons on the window must function as specified in the 6.1.1 |  |  |  |  |  |
| ## Generate MKT |  |  |  |  |  |  |  |  |  |  |
| a) | Click On Generate MKT in the Data Log option |  |  | The Generate MKT window will open |  |  |  |  |  |  |
| b) | Enter the Filter Option Click Print |  |  | Click on Generate MKT |  |  |  |  |  |  |
| c) | Authentication window open |  |  | Authentication Window must open and Entering the login details with comments |  |  |  |  |  |  |
| d) | Click Log In |  |  | MKT Should Generating and displayed in MKT Analysis |  |  |  |  |  |  |
| ## MKT Analysis |  |  |  |  |  |  |  |  |  |  |
| a) |  | Click On MKT Analysis in the Data Log option | The MKT Analysis window will open |  |  |  |  |  |  |  |
| b) |  | Enter the Filter Option Click Print | Print Preview of filtered data in PDF opens |  |  |  |  |  |  |  |
| c) |  | Click On Print | Should print the displayed MKT for the selected sensor |  |  |  |  |  |  |  |
| d) |  | Buttons | The Available buttons on the window must function as specified in the 6.1.1 |  |  |  |  |  |  |  |
| Sr. No. |  | Technical data Version |  |  |  |  |  | Observation | Test result met (Yes/No) | Results and Comments |
| ## Alarm Log |  |  |  |  |  |  |  |  |  |  |
|  |  | Click On the Alarm Log option in the Data Log option |  |  |  | The Alarm Log window must Opens |  |  |  |  |
|  |  | Enter the Filter Option |  |  |  | The Alarm logs for the filtered option must get populate. |  |  |  |  |
|  |  | Print |  |  |  | Should open the Alarm print preview window |  |  |  |  |
|  |  | Print Button |  |  |  | The displayed alarms must get printed by clicking on the print command. |  |  |  |  |
|  |  | Select the check box in Alarm Log window to select the alarms and Click On Acknowledge |  |  |  | Login Window must open for authentication. |  |  |  |  |
|  |  | Login Window Click Log In |  |  |  | Entering the login details with comments acknowledge the selected alarms and the details moves from the Alarm Log window to Alarm Audit Trail. |  |  |  |  |

| Sr. No. | Technical data Version |  | Observation | Test result met (Yes/No) | Results and Comments |
| --- | --- | --- | --- | --- | --- |
| # Audit Trail |  |  |  |  |  |
| ## Alarm Audit Trail |  |  |  |  |  |
|  | Click On the Alarm Audit Trail option in the Audit Trail | The Alarm Audit Trail window Opens |  |  |  |
|  | Enter the Filter Option Click Print | The Audit Trail for the filtered details must populate. |  |  |  |
|  | Print Report | Print Report command Should open the Alarm Audit Trail print preview window |  |  |  |
|  | Print Format | Select the different option in filter option to view reports in different format |  |  |  |
|  | Buttons | Check Print Report Button Function |  |  |  |
| ## Instrument Audit Trail |  |  |  |  |  |
|  | Click On the Instrument Audit Trail option in the Audit Trail | The Instrument Audit Trail window Opens |  |  |  |

| Sr. No. | Technical data Version |  | Observation | Test result met (Yes/No) | Results and Comments |
| --- | --- | --- | --- | --- | --- |
|  | Enter the Filter Option Click Print | The Audit Trail for the filtered details must populate. |  |  |  |
|  | Print | Print Report command Should open the Audit Trail print preview window |  |  |  |
|  | Print Format | Select the different option in filter option to view reports in different format |  |  |  |
|  | Buttons | Check Print Report Button Function |  |  |  |
| ## User Audit Trail |  |  |  |  |  |
|  | Click On the User Audit Trail option in the Audit Trail | The User Audit Trail window Opens |  |  |  |
|  | Enter the Filter Option Click Print | The Audit Trail for the filtered details must populate. |  |  |  |
|  | Print | Print command Should open the User Audit Trail print preview window |  |  |  |
|  | Print Format | Select the different option in filter option to view reports in different format |  |  |  |
| Sr. No. | Technical data Version |  | Observation | Test result met (Yes/No) | Results and Comments |
|  | Buttons | Check Print Report Button Function |  |  |  |
| ## Email Audit Trails (Only if Email Alert Is Active) |  |  |  |  |  |
|  | Click On the Email Audit Trail option in the Audit Trail | The Email Audit Trail window Opens |  |  |  |
|  | Enter the Filter Option Click Print | The Audit Trail for the filtered details must populate |  |  |  |
|  | Print Report | Print command Should open the Audit Trail print preview window |  |  |  |
|  | Print Format | Select the different option in filter option to view reports in different format |  |  |  |
|  | Buttons | Check Print Report Button Function |  |  |  |
| ## SMS Audit Trails (If SMS is Active) |  |  |  |  |  |
|  | Click On the SMS Audit Trail option in the Audit Trail | The SMS Audit Trail window Opens |  |  |  |
|  | Enter the Filter Option Click Print | The Audit Trail for the filtered details must populate and must print the displayed audit trail. |  |  |  |

| Sr. No. | Technical data Version |  | Observation | Test result met (Yes/No) | Results and Comments |
| --- | --- | --- | --- | --- | --- |
|  | Print Report | Print command Should open the Audit Trail print preview window |  |  |  |
|  | Print Format | Select the different option in filter option to view reports in different format |  |  |  |
|  | Buttons | Check Print Report Button Function |  |  |  |
| ## Reviewed/ Approved Audit Trails |  |  |  |  |  |
|  | Click On the Reviewed/Approved Audit Trail option in the Audit Trail | The Reviewed/Approved Audit Trail window Opens |  |  |  |
|  | Enter the Filter Option Click Print | The Audit Trail for the filtered details must populate and must print the displayed audit trail. |  |  |  |
|  | Print | Print command Should open the Audit Trail print preview window |  |  |  |
|  | Print Format | Select the different option in filter option to view reports in different format |  |  |  |
|  | Buttons | Check Print Report Button Function |  |  |  |

| Sr. No. | Technical data Version |  | Observation | Test result met (Yes/No) | Results and Comments |
| --- | --- | --- | --- | --- | --- |
| # Archive Analysis |  |  |  |  |  |
| ## Archive Analysis: - For archive analysis restore the backup main data base data i.e. NECLOG_3_0.bak file using data base restore utility. Select The validation Option for restoring the data. Once the features are tested restore the Archive back up NECLOG3_0_Archive.bak file again so as to avoid the mismatch in future. Note: Before perform above activity take backup of Neclog 3.0 databases & Keep databases on other folder for security. |  |  |  |  |  |
| ### Database Restore Start the Database Restore Utility in the restore utility folder and follow the steps for Database restore |  |  |  |  |  |
| a) | Start the Database Restore Utility | The Restore utility must start with Login Screen |  |  |  |
| b) | SQL Login Details | Enter the Valid SQL SERVER Login Name and Password to open the Restore utility |  |  |  |
| c) | Select the “Restore Blank Database” Option to check the possibility of restoration the data backup” | The Error Message must get Populate |  |  |  |
| d) | Select the “Already Existing Database” Option to restore the data backup” | The Data restore Window must get opened |  |  |  |

| Sr. No. | Technical data Version |  |  | Observation | Test result met (Yes/No) | Results and Comments |
| --- | --- | --- | --- | --- | --- | --- |
| e) | The below option must be available - Select Path - Data Restore - Close - Main Database - Archive Database - Validation | Select The option to restore the Backup |  |  |  |  |
| f) | Click On Select Path to restore the Backup | The Browse window must populate to select the Backup |  |  |  |  |
| g) | Navigate the appropriate folder to select the wrong backup and click on Ok | The selected backup must get displayed in the utility window |  |  |  |  |
| h) | Click on Main Database and click on restore | Error message must populate |  |  |  |  |
| i) | Click on Archive Database and click on restore | Error message must populate |  |  |  |  |
| j) | Click on Validation and click on restore | Error message must populate |  |  |  |  |
| k) | Navigate the appropriate folder to select the Correct NECLOG_3_0.bak backup and click on Ok | The selected backup must get displayed in the utility window |  |  |  |  |
| l) | Click on Main Database and click on restore | The selected file must get restored successfully |  |  |  |  |
| Sr. No. | Technical data Version |  |  | Observation | Test result met (Yes/No) | Results and Comments |
| m) | Check the data in Main Database for confirmation | The Files along with data must be correct |  |  |  |  |
| n) | Navigate the appropriate folder to select the Correct archive NECLOG_3_0_Archive.bak backup and click on Ok | The selected backup must get displayed in the utility window |  |  |  |  |
| o) | Click on Archive Database and click on restore | The selected file must get restored successfully |  |  |  |  |
| p) | Check the data in Archive Database for confirmation | The Files along with data must be correct |  |  |  |  |
| q) | Navigate the appropriate folder to select the Correct NECLOG_3_0.bak backup and click on Ok | The selected backup must get displayed in the utility window |  |  |  |  |
| r) | Click on Validation and click on restore | The selected file must get restored successfully |  |  |  |  |
| s) | Check the data in Archive Database for confirmation | The Files along with data must be correct |  |  |  |  |
| Sr. No. | Technical data Version |  | Observation |  | Test result met (Yes/No) | Results and Comments |
| ## Data Log Analysis |  |  |  |  |  |  |
| ### View Data Log |  |  |  |  |  |  |
| a) | Click On the View Data Log in Data Log Analysis | The View Data log window will open |  |  |  |  |
| b) | Enter the Filter Option Click Print | Print Preview of filtered data in PDF opens |  |  |  |  |
| c) | Click On Graph | The Graph for the selected filtered details must get displayed |  |  |  |  |
| e) | Buttons | Check Print Report Button Function |  |  |  |  |
| ### View Reviewed Log |  |  |  |  |  |  |
| a) | Click On the View Reviewed Data Log in Data Log Analysis | The View Reviewed Data log window will open |  |  |  |  |
| b) | Enter the Filter Option Click Print | Print Preview of filtered data in PDF opens |  |  |  |  |
| Sr. No. | Technical data Version |  | Observation |  | Test result met (Yes/No) | Results and Comments |
| c) | Click On Graph | The Graph for the selected filtered details must get displayed |  |  |  |  |
| d) | Buttons | Check Print Report Button Function |  |  |  |  |
| ### View Approved Data |  |  |  |  |  |  |
| a) | Click On the View Approved Data Log in Data Log Analysis | The View Approved Data log window will open |  |  |  |  |
| b) | Enter the Filter Option Click Print | Print Preview of filtered data in PDF opens |  |  |  |  |
| c) | Click On Graph | The Graph for the selected filtered details must get displayed |  |  |  |  |
| d) | Buttons | Check Print Report Button Function |  |  |  |  |
| ### MKT Analysis |  |  |  |  |  |  |
| a) | Click On MKT Analysis in the Data Log option | The MKT Analysis window will open |  |  |  |  |
| b) | Enter the Filter Option Click Print | Print Preview of filtered data in PDF opens |  |  |  |  |
| c) | Click On Print | Should print the displayed MKT for the selected sensor |  |  |  |  |
| Sr. No. | Technical data Version |  | Observation |  | Test result met (Yes/No) | Results and Comments |
| d) | Buttons | Check Print Report Button Function |  |  |  |  |
| ### Alarm Log |  |  |  |  |  |  |
| a) | Click On the Alarm Log option in the Audit Trail window | The Alarm Log window Opens |  |  |  |  |
| b) | Enter the Filter Option | The Alarm logs for the filtered option must populate. |  |  |  |  |
| c) | Print | Should open the Alarm print preview window |  |  |  |  |
| d) | Print Button | The displayed alarms must get printed by clicking on the print command. |  |  |  |  |
| e) | Buttons | Check Print Report Button Function |  |  |  |  |
| ## Audit Trail |  |  |  |  |  |  |
| a) | Alarm Audit trail | Alarm Audit Trail window must open |  |  |  |  |
| b) | Instrument Audit trail | Instrument Audit Trail window must open |  |  |  |  |
| c) | User Audit Trail | User Audit Trail window must open |  |  |  |  |
| Sr. No. | Technical data Version |  | Observation |  | Test result met (Yes/No) | Results and Comments |
| d) | Email Audit Trail | Email Audit Trail window must open |  |  |  |  |
| e) | SMS Audit Trail | SMS Audit Trail window must open |  |  |  |  |
| f) | Reviewed/Approved Audit Trail | Reviewed/Approved Audit Trail window must open |  |  |  |  |
| ## Data Restore Note: Once the Archive Test over Restore the Archive Data base again as defined above in OQ Test Case No 14.1.1 |  |  |  |  |  |  |
| a) | Navigate the appropriate folder to select the Correct archive NECLOG_3_0_Archive.bak backup and click on Ok | The selected backup must get displayed in the utility window |  |  |  |  |
| b) | Click on Archive Database and click on restore | The selected file must get restored successfully |  |  |  |  |
| c) | Check the data in Archive Database for confirmation | The Files along with data must be correct |  |  |  |  |
| # Company Information |  |  |  |  |  |  |
|  | Company Information | This window gives the detail description about the company profile |  |  |  |  |

| Sr. No. | Test |  | Observation | Test result met Yes/No | Results and Comments |
| --- | --- | --- | --- | --- | --- |
| # User Access Rights Groups Verification |  |  |  |  |  |
| a) | Click on User Group option in Administrator window | The User Group window will open and must display the added user groups |  |  |  |
| b) | To add/edit new User group click on Add/edit and enter Name Details Select Access Rights Click on Save | Details Entered/edited |  |  |  |
| c) | Create 3 different User Groups | User Groups with defined rights must get created |  |  |  |
| d) | 1st User Groups Created with rights | Define the rights for User Group ____________ |  |  |  |
| e) | 2nd User Groups Created with rights | Define the rights for User Group ____________ |  |  |  |
| f) | 3rd User Groups Created with rights | Define the rights for User Group ____________ |  |  |  |

| Sr. No. | Test |  | Description |  | Expected Result |  | Observation | Test result met Yes/No | Results and Comments |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| # Password Integrity---Admin User |  |  |  |  |  |  |  |  |  |
| Description |  | Verify the facilities of password change, auto password expiry, minimum password length, user name & password uniqueness, login/logout, account lockout facilities, user creation/deletion options in Software |  |  |  |  |  |  |  |
| ## Login name Uniqueness / Password Complexity |  |  |  |  |  |  |  |  |  |
| a) | Password –Minimum Length |  |  | Check Password character length |  | Password length should be 8-15 characters. |  |  |  |
| b) | Login name Uniqueness |  |  | Check Login uniqueness |  | System should not allow the generation of same users with unique login name. |  |  |  |
| c) | Password Complexity |  |  | Check password complexity |  | Password combination should complex. |  |  |  |
| d) | Account lock |  |  | Check wrong password up to …….. times as set in application setup |  | Username account should be blocked |  |  |  |
| ## Password Integrity |  |  |  |  |  |  |  |  |  |
| a) | Combination |  |  | No Login name No Password | Login must not be allowed |  |  |  |  |
| b) | Combination |  |  | Wrong Login name Wrong Password | Login must not be allowed |  |  |  |  |
| c) | Combination |  |  | Right Login name Wrong Password | Login must not be allowed |  |  |  |  |
| d) | Combination |  |  | Wrong Login name Right Password | Login must not be allowed |  |  |  |  |
| e) | Combination |  |  | Right Login name and Right Password | Login must be allowed |  |  |  |  |

| Sr. No. | Test |  | Description |  | Expected Result | Observation | Test result met Yes/No | Results and Comments |
| --- | --- | --- | --- | --- | --- | --- | --- | --- |
| # Password Integrity—Application User |  |  |  |  |  |  |  |  |
| Description |  | Verify the facilities of password change, auto password expiry, minimum password length, user name & password uniqueness, login/logout, account lockout facilities, user creation/deletion options in Software |  |  |  |  |  |  |
| ## Login name Uniqueness / Password Complexity |  |  |  |  |  |  |  |  |
| a) | Password –Minimum Length |  |  | Check Password character length | Password length should be 8-15 characters. |  |  |  |
| b) | Login name Uniqueness |  |  | Check Login uniqueness | System should not allow the generation of same users with unique login name. |  |  |  |
| c) | Password Complexity |  |  | Check password complexity | Password combination should complex. |  |  |  |
| d) | Account lock |  |  | Check wrong password up to ….. times as set in application Setup | Username account should be blocked |  |  |  |
| ## Password Integrity |  |  |  |  |  |  |  |  |
| a) | Combination |  |  | No Login name No Password | Login must not be allowed |  |  |  |
| b) | Combination |  |  | Wrong Login name Wrong Password | Login must not be allowed |  |  |  |
| c) | Combination |  |  | Right Login name Wrong Password | Login must not be allowed |  |  |  |
| d) | Combination |  |  | Wrong Login name Right Password | Login must not be allowed |  |  |  |
| e) | Combination |  |  | Right Login name and Right Password | Login must be allowed |  |  |  |

| Sr. No. | Test | Description | Expected Result | Observation | Test result met Yes/No | Results and Comments |
| --- | --- | --- | --- | --- | --- | --- |
| # Admin User Addition / Deactivation |  |  |  |  |  |  |
| a) | User Creation | Click on Admin User to create a user | System should allow creation of user_____________ with defined rights |  |  |  |
| b) | User Deletion if User not logged in anytime | Go to Admin User, Select the user and click on delete | System should allow deletion of new user |  |  |  |
| c) | User Deletion if User has logged in at least once | Go to Admin User, Select the user and click on delete | System should not allow deletion of user if the user has not logged in previously |  |  |  |
| d) | User Deactivation | Go to Admin User, double click on the user and change status from active to deactivate | System should allow Deactivation of user |  |  |  |
| e) | User Rights Modification | Go to Admin User, double click on the user and change group to new group | System should allow modification of user rights. |  |  |  |
| f) | Compulsory Password Change | Login to software with new user created | The software must ask for compulsory password change during first login |  |  |  |
| g) | User Password Expiry | Set the password expiry duration as 1 day in Application Setup and check for password change popup after expiry duration | The change password window must open when password duration expires |  |  |  |

| Sr. No. | Test | Description | Expected Result | Observation | Test result met Yes/No | Results and Comments |
| --- | --- | --- | --- | --- | --- | --- |
| # Application User Addition / Deactivation |  |  |  |  |  |  |
| a) | User Creation | Click on Application User to create a user | System should allow creation of user_____________ with defined rights |  |  |  |
| b) | User Deletion if User not logged in anytime | Go to Application User, Select the user and click on delete | System should allow deletion of new user |  |  |  |
| c) | User Deletion if User has logged in at least once | Go to Application User, Select the user and click on delete | System should not allow deletion of user if the user has not logged in previously |  |  |  |
| d) | User Deactivation | Go to Application User, double click on the user and change status from active to deactivate | System should allow Deactivation of user |  |  |  |
| e) | User Rights Modification | Go to Application User, double click on the user and change group to new group | System should allow modification of user rights. |  |  |  |
| f) | Compulsory Password Change | Login to software with new user created | The software must ask for compulsory password change during first login |  |  |  |
| g) | User Password Expiry | Set the password expiry duration as 1 day in Application Setup and check for password change popup after expiry duration | The change password window must open when password duration expires |  |  |  |

| Sr. No. | Test |  | Observation |  |  | Test result met Yes/No | Results and Comments |
| --- | --- | --- | --- | --- | --- | --- | --- |
| # Admin User Access Rights Verification |  |  |  |  |  |  |  |
| a) | Go to Admin User Option in Administrator | The Admin User window must open |  |  |  |  |  |
| b) | Use 3 different user with different rights | The User with defined rights must get created | No of User | Name | User Group |  |  |
|  |  |  | 1 |  |  |  |  |
|  |  |  | 2 |  |  |  |  |
|  |  |  | 3 |  |  |  |  |
| c) | Login to software with User 2 details | The software must give assess to the software for User 1________________ with defined user rights of ___________ |  |  |  |  |  |
| d) | Login to software with User 1 details | The software must give assess to the software for User 2________________ with defined user rights of ___________ |  |  |  |  |  |
| e) | Login to software with User 3 details | The software must give assess to the software for User 3________________ with defined user rights of ___________ |  |  |  |  |  |
| Sr. No. | Test |  | Observation |  |  | Test result met Yes/No | Results and Comments |
| # Application User Access Rights Verification |  |  |  |  |  |  |  |
| a) | Go to Application User Option in Configuration | The Application User window must open |  |  |  |  |  |
| b) | Use 3 different user with different rights | The User with defined rights must get created | No of User | Name | User Group |  |  |
|  |  |  | 1 |  |  |  |  |
|  |  |  | 2 |  |  |  |  |
|  |  |  | 3 |  |  |  |  |
| c) | Login to software with User 2 details | The software must give assess to the software for User 1________________ with defined user rights of ___________ |  |  |  |  |  |
| d) | Login to software with User 1 details | The software must give assess to the software for User 2________________ with defined user rights of ___________ |  |  |  |  |  |
| e) | Login to software with User 3 details | The software must give assess to the software for User 3________________ with defined user rights of ___________ |  |  |  |  |  |

| Sr. No. | Test |  | Observation |  | Test result met Yes/No | Results and Comments |
| --- | --- | --- | --- | --- | --- | --- |
| # Software Functionality testing |  |  |  |  |  |  |
| ## Power Failure Condition |  |  |  |  |  |  |
| Description |  | The software should not be affected by power failure and should start automatically as per user settings after power has resumed. |  |  |  |  |
| a) | Procedure | Turn OFF the computer system power for 30 seconds and restart: - Does the computer system start normally? - Does the software start normally? |  |  |  |  |
| ## Auto Backup |  |  |  |  |  |  |
| a) | Check the Backup path for auto backup file | Backup namely - NECLOG_3_0.bak - NECLOG_3_0_Archive.bak Must be there at defined time |  |  |  |  |
| ## Auto Logout |  |  |  |  |  |  |
| a) | Login to Software | Home screen must open |  |  |  |  |
| b) | Keep the software and computer idle for 10 mins | Logon Screen must get open if the system is idle for more than 10 mins |  |  |  |  |
| ## Auto Password Save |  |  |  |  |  |  |
| Description |  | The software /browser must not popup any window for saving the password |  |  |  |  |
| a) | Login to Software | Does the software/browser ask for password save option when logging in software? |  |  |  |  |

| Sr. No. | Technical data Version |  |  |  |  |  | Test result met (Yes/No) | Results and Comments |
| --- | --- | --- | --- | --- | --- | --- | --- | --- |
|  | Audit Trail For |  | Audit Trail Occurs After Record Modification? (Y/N/NA) | New & Old Data Stored In Audit Trail? (Y/N/NA) | Audit Trail Data Stored With Timestamp? (Y/N/NA) | Audit Trail Data Stored With Username? (Y/N/NA) |  |  |
| # Audit trail Functionality testing |  |  |  |  |  |  |  |  |
| Description |  |  | The verify the software generates audit trails of any changes made from the software |  |  |  |  |  |
| a) |  | Alarm Audit Trail |  |  |  |  |  |  |
| b) |  | Instrument Audit Trial |  |  |  |  |  |  |
| c) |  | User Audit Trial |  |  |  |  |  |  |
| d) |  | Email Audit Trail |  |  |  |  |  |  |
| e) |  | SMS Audit Trail |  |  |  |  |  |  |
| f) |  | Reviewed/ Approved Data log Audit trial |  |  |  |  |  |  |

# References / Glossary

| Abbreviation | Explanation |
| --- | --- |
| OQ | Operation Qualification |
| GUI | Graphical User Interference |
| LAN | Local Area Network |
| Comm. | Communication |
| PID | Proportion Integral Derivative control action |
| SMTP | Simple Mail Transfer Protocol |
| Email ID | Electronic Mail ID |
| MKT | Mean Kinetic Temperature |
| PDF | Portable Document Form |
| SMS | Short Message Service |
| IP | Internet protocol |
| LUX | Luminous Flux |
| UV | Ultra Violet |
| RTC | Real Time Clock |
| Config | Configuration |
| IO | Input /Output |

# Deviation Report

Report Number: ………………………………………………….

Deviations to the acceptance criteria that occurred during testing need to be stated in the deviation report. One deviation report must be created for each deviation and the report number to be indicated accordingly.

The deviation report can be copied as required.

| Description of problem | ……………………………………………………………………………………… ……………………………………………………………………………………… ……………………………………………………………………………………… ……………………………………………………………………………………… ……………………………………………………………………………………… ……………………………………………………………………………………… |
| --- | --- |
| Measures for trouble shooting | ……………………………………………………………………………………… ……………………………………………………………………………………… ……………………………………………………………………………………… ……………………………………………………………………………………… ……………………………………………………………………………………… ……………………………………………………………………………………… ……………………………………………………………………………………… |
| Remarks | ……………………………………………………………………………………… ……………………………………………………………………………………… ……………………………………………………………………………………… |
| Executed By Verified/ accepted by | Name:………………………Date: . . . . . . . . . . . . . . . . Signature: . . . . .. . . . . . . . . . . . . . . Name:………………………Date: . . . . . . . . . . . . . . . . Signature: . . . . .. . . . . . . . . . . . . . |

# Appendix

| Annexure Number | Contents | Number of pages |
| --- | --- | --- |

| Annexure Number | Contents | Number of pages |
| --- | --- | --- |

| Annexure Number | Contents | Number of pages |
| --- | --- | --- |

# Conclusion

. . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . .  . . . . . . . . . . . . . . . . . . . . . . . . . .

. . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . .. . . . . . . . . . . . . . . . . . . . . . . . . . . .

. . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . .  . . . . . . . . . . . . . . . . . . . . . . . . . . .

. . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . .  . . . . . . . . . . . . . . . . . . . . . . . . . . .

. . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . .. . . . . . . . . . . . . . . . . . . . . . . . . . . .

. . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . .  . . . . . . . . . . . . . . . . . . . . . . . . . . .

. . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . .. . . . . . . . . . . . . . . . . . . . . . . . . . . .

. . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . .. . . . . . . . . . . . . . . . . . . . . . . . . . . .

. . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . .. . . . . . . . . . . . . . . . . . . . . . . . . . . .

. . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . .. . . . . . . . . . . . . . . . . . . . . . . . . . . .

. . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . .. . . . . . . . . . . . . . . . . . . . . . . . . . .

# Approval of test results

The Operational Qualification is approved if all acceptance criteria are fulfilled and the notations / deficiencies registered in the deviation protocol are corrected or fully accepted.

Approved By:

| Tests executed | Name | Signature and Date |
| --- | --- | --- |
| Executed By | . . . . . . . . . . . . . . . . . . . | . . . . . . . . . . . . . . . . . . |
| All Test criteria have been fulfilled | Name | Signature and Date |
| Verified By | . . . . . . . . . . . . . . . . . . . | . . . . . . . . . . . . . . . . . . |
| All Test criteria have been fulfilled | Name | Signature and Date |
| Approved By | . . . . . . . . . . . . . . . . . . . | . . . . . . . . . . . . . . . . . . |